| Internal (confidential) | VS: | MAH | 25.02.2016 |
| --- | --- | --- | --- |

| Unit: | AT-104 PC | Software: | SDS-104 |
| --- | --- | --- | --- |

| Software | Name | Version: from |  | to |
| --- | --- | --- | --- | --- |
| SDS-104 | Install.exe |  | 2.71.0 | 2.72.0 |

| Languages: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Tested by | GER | ENG | USA | FRA | ITA | SPA | POR | SWE | FIN | NLA | RUS | NOR | CZE | SER |
| SCHILLER | X | X |  |  |  |  |  |  |  |  |  |  |  |  |
| Distributor |  |  | X | X | X | X | X | X |  | X | X | X | X | X |
| Tested by | DAN | GRE | POL | SLV | CRO | TUR | ENG(ZA) | HUN | SLK |  |  |  |  |  |
| SCHILLER |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Distributor | X | X | X | X | X | X | X | X | X |  |  |  |  |  |

| Compatibility tested with: | Version | Comments |
| --- | --- | --- |
| AT-104 PC | 1.86 |  |
| SCS | 2.6.4 |  |
| SemaImport | 2.72.0 |  |
| ADS | 11.10 |  |
| SCHILLER Server | 3.7.0 / 3.8.0 |  |
| DARWIN | 2.6.0 |  |
| SEMA-200 | 2.72.0 |  |

| System requirements |  |
| --- | --- |
| Hardware: | 1.0 GHz 1024 MB RAM (Free Memory) |
| Platform: | Windows 7 32 bit Windows 7 64 bit Windows 8.1 64 bit Windows 10 64 bit |

General

New functions

Send interpretation and measurements with a PDF Upload to SEMA3

Show dialog on export if recording has been already exported (only if SEMA3 integration is set)

Changed functions

#19429:	Other Export formats are possible to use if SEMA3 Integration is enabled.

#19414: 	Export Interpretation and global measurement values on PDF-SEMA3 Integration

#19422: 	Show Dialog on Export if recording has already been exported (SEMA3 Integration)

Error corrections

#18947: 	Recordings can be sent duplicated to SEMA3

#18127/18126: Path to PDF file in GDT 6305 required, not only the filename

#19416: 	Rhythm and Exercise ECG’s in Cabrera cannot be uploaded to SEMA3

#17914: 	Unknown Serial Number, Device ID and wrong Model in device list

#17965: 	PC-Spiro export to SEMA3 fails due missing SN in XML

#19325: 	SDS-104 Calibration Summary

Known errors

none

In future versions

n/a

Changes in the user guide

none

Changed HW components

n/a